## Slide 1: DNA Puff qPCR
DATE
Project/plate #:
e.g. Validation Practice Plate 001
Your Name
Fill in the primer names on next slide of primers you will use and columns they will take up.

## Slide 2
1:1
Validation Plate: DATE
1:4 or 1:10
1:16 or 1:100
1:64 or 1:1000
Add primer pair master mixes column by column
Add gDNA dilution series concentrations row by row
Seal with lid, spin, bring to qPCR machine
gDNA dilution
Each plate should have the normalizer site as well to look at relative efficiency.
JU-control-1
MSF-target-1
MSF-control-1
JU-control-2
JU-control-3
JU-control-4
JU-control-5
JU-target-1
JU-target-2
JU-target-3
JU-target-4
JU-target-5

## Slide 3
Validation plate protocol - 15 ul Reaction
Sign up for the qPCR machine for when you expect to be done preparing the plate.
Over-estimating is better than under-estimating, you can always store the taped-up plate wrapped in tin foil (to protect from light) at 4*C for a while. Most important is making sure you get a time-slot at all. Wipe down entire area with 70% etoh. Get rid of dust from shelves. Use flame. Get 3 ice buckets.
Step 1 – if have not made 1 uM primer solutions, do so:
- Add 180 ul of UPW to 12 pre-labeled tubes
- Add 20 ul of corresponding 10 uM primer to each
- Thaw primers on ice, and set up workbench (fill 3 ice buckets, empty waste container, wipe down surfaces with EtOH, etc.)
Step 2 -- gDNA dilutions
For Sciara (can also do 10-fold dilutions as with yeast – but may want to revisit the Upper/Lower limits of series): Starting concentration is 1 ng/ul and will use 2 ul per reaction (for 1:1). Qubit gDNA first to get concentration.
Then dilute down to 1 ng/ul in sufficient volume for all reactions (24/plate) and dilution series.
This is 1:1. Need at least 48 ul at 2 ul per reaction, 24 rxn per plate + 25 ul for dilution + pipeting error. So need >75 ul – get 100 ul at 1 ng/ul (100 ng total is all you need per 96-well validation plate).
Make 1:4 by combining 25 ul 1:1 with 75 ul UPW. [[Use new tips!]]
Make 1:16 by combining 25 ul 1:4 with 75 ul UPW. [[Use new tips!]]
Make 1:64 by combining 25 ul 1:16 with 75 ul UPW. [[Use new tips!]]
For yeast:
Starting concentration is 0.5 ng/ul and will use 2 ul per reaction. Qubit gDNA first to get concentration.
Then dilute down to 1 ng/ul in sufficient volume for all reactions (24/plate) and dilution series.
This is 1:1. Need at least 48 ul at 2 ul per reaction, 24 rxn per plate + 10 ul for dilution + pipeting error. So need >60 ul – get 70 ul at 0.5 ng/ul (35 ng total is all you need per 96-well validation plate).
Make 1:10 by combining 10 ul 1:1 with 90 ul UPW. [[Use new tips!]]
Make 1:100 by combining 10 ul 1:10 with 90 ul UPW. [[Use new tips!]]
Make 1:1000 by combining 10 ul 1:100 with 90 ul UPW. [[Use new tips!]]
Notes: For Human:
Starting concentration is 0.5 ng/ul and will use 2 ul per reaction. Qubit gDNA first to get concentration.
Then dilute down to 1 ng/ul in sufficient volume for all reactions (24/plate) and dilution series.
This is 1:1. Need at least 48 ul at 2 ul per reaction, 24 rxn per plate + 10 ul for dilution + pipeting error. So need >60 ul – get 70 ul at 0.5 ng/ul (35 ng total is all you need per 96-well validation plate).
Make 1:10 by combining 10 ul 1:1 with 90 ul UPW. [[Use new tips!]]
Make 1:100 by combining 10 ul 1:10 with 90 ul UPW. [[Use new tips!]]
Make 1:1000 by combining 10 ul 1:100 with 90 ul UPW. [[Use new tips!]]

## Slide 4
Validation plate protocol - 15 ul Reaction
Step 3 – make all master mixes – keep on ice – 8 rxns per primer pair; 12 primer pairs per plate; Make enough for 9 rxn
- label 12 tubes with primer pair names
Tape pipette to appropriate volume
in each add 
After all in,
Short spin
mix up/down twice
Short spin down
(put primers back in freezer)
Step 4 - load columns of plate with corresponding primer master mixes:
Place plate on empty tip platform, and place platform on ice with barcode facing you
13 ul of master mix in each well
Master mix ‘n’ goes into every row cell of column ‘n’ (see plate guide)
---> keep all columns of plate to right of current column covered with optical tape/lid (without exposing sticky part)
--> Only move pipette over the covered regions and current well
--> Drop solution along side of well, using same tip for each column
(need to change tips when going to next column for new master mix)
(in step 5 you will need to change tips for every single well)
Step 5 - load rows with corresponding DNA concentrations (high to low):
Add 2ul of DNA to each well. Drop DNA along side of each well. Often the drop won’t go down the well. Consider this helpful as it indicates you already put DNA in the well. Do not try to force it down.
Whereas primers were added left to right, one column at a time, DNA is now added top to bottom, one row at a time (or 2 rows at a time depending on how you think/work).
Rows A,B = 1:1 (1:1)
Rows C,D = 1:10 (1:4) or (1:10)
Rows E,F = 1:100 (1:16) or (1:100)
Rows G,H = 1:1000 (1:64) or (1:1000)
9 reactions using 1 uM:
67.5 ul 2X SYBR Green
22.5 ul UPW
13.5 ul 1 uM Fwd
13.5 ul 1 uM Rev
117 ul

## Slide 5
Validation plate protocol - 15 ul Reaction
Step 6 - process plate
- add seal
  spin down at 1000 RPM for 1 min
  Will need to change rotor to the plate-holding rotor
  Check plate after spinning down to make sure no bubbles are present  Spin plate at 1500 RPM to pop any remaining bubbles. Can also flick to help pop bubbles before spinning again.
Step 7 . Load plate into qPCR machine, set it up on computer, run it
It is thought that one should restart both the qPCR machine and computer before beginning (no password needed, just press “OK”).
Click/open the qPCR program on computer
New document
Accept settings, click next (validation plate  absolute quant)
Add appropriate primer sets or make them. Remember to choose unique color if making. Reporter Dye = SYBR. Quencher Dye = None. Passive Reference Dye = ROX.
Click Next
Highlight appropriate wells and label as that particular primer set
Within each primer set, label wells as either standard, NTC, or “unknown” according to what they are
Click Next or Finish
Be in Instrument tab
Add dissociation stage
make sample volume = SampleVol (i.e 15-25 ul)
Take plate from ice, wipe bottom of wells with kimwipe, put plate into machine
start the qPCR cycles
Notes: Add 2 ul DNA one row at a time according to plate guide (keeping all rows beneath it covered)
Spin down ~1-1.5k RPM for 1 min

## Slide 6
Validation plate protocol - 15 ul Reaction
Step 8. Finalize qPCR run and obtain data
When qPCR complete, go to Analysis tab
start cycle = 3
double click y-axis, make linear
look at cycle number where traces begin to lift off baseline
subtract 3 from that cycle number and make that the “end cycle”
This is setting cycle threshold
Export results to thumb drive
Step 9. Analyze results in Excel or R
– results – check slope, eff, Rsqr – pass/fail – example:
primerPair slope efficiency Rsqr pass
1 n1 -3.120475 1.0915297 0.9934686 1
2 n2 -3.711285 0.8597238 0.9795718 0
3 n3 -3.229430 1.0401033 0.9985078 1
4 n4 -3.254095 1.0291077 0.9960876 1
5 n5 -3.218725 1.0449468 0.9955501 1
6 n6 -3.155280 1.0745747 0.9966649 1
7 n7 -3.246655 1.0324006 0.9964864 1
8 n8 -3.253750 1.0292600 0.9846853 1
9 n9 -3.194015 1.0562957 0.9970389 1
10 n10 -3.145015 1.0795219 0.9932453 1
11 n11 -3.192585 1.0569598 0.9935030 1
12 n12 -3.352260 0.9874958 0.9971362 1
Relative Eficiency
Using a given normalizer --- this helps pick the best normalization site, if testing more than 1. The goal is to have the same ratio of test:normalizer across serial dilution conditions.In other words, a line drawn through should have a slope close to 0.
Example Norm-site-1 as normalizer:
primer slope pass
n1 0.0735400000000006 TRUE
n2 -0.517269999999997 FALSE
n3 -0.0354149999999997 TRUE
n4 -0.0600799999999992 TRUE
n5 -0.0247099999999982 TRUE
n6 0.0387350000000001 TRUE
n7 -0.0526399999999992 TRUE
n8 -0.0597349999999995 TRUE
n9 0 TRUE
n10 0.0490000000000006 TRUE
n11 0.00143000000000023 TRUE
n12 -0.158244999999999 FALSE
When qPCR finishes, take notes on melt curve and what each reaction looks like.
For those that do not immediately pass validation, check if there is a clear outlier. When that is removed, does it then pass? If so, it is worth trying re-validation.

## Slide 7: Dissociation notes
Sample/condition | observations

## Slide 8: Amp plot notes
Sample/condition | spacing | tightness | Not tight | 

## Slide 9: Primer pair test restults
Primer pair | Slope | Efficiency | Rsqr | pass | date

## Slide 10: Primer pair vs normalizer relative efficiency test results
Normalizer pair | Test pair | Slope | Pass | date